NEW HORIZONS DIAGNOSTICS

SMART II Salmonella

Test Kit (20 cassettes)

Cat. #:  89-127001-C

INTENDED USE

NHD SMART II Salmonella test is a qualitative rapid test for the detection of Salmonella in enriched media.  This test is for point-of-care use or reference lab settings.  The test is designed as an aid in the detection of Salmonella and should be used as an adjunct to culture.

INTRODUCTION

Most foods harbor microorganisms in differing concentrations.  Salmonella species are one of the major causes of foodborne illness worldwide.  There are over 1500 Salmonella serotypes (species).  Many serotypes are capable of causing salmonellosis in humans; however, only a small number of these serotypes have been associated with foodborne illness.  S.typhimurium and S enteritidis have been implicated in most of foodborne outbreaks.

Symptoms appear 8-42 hours after ingestion of Salmonella-contaminated food.  Symptoms include diarrhea, vomiting, chills, fever, nausea, abdominal cramps, and prostration.

In order to detect low levels of Salmonella in foods, a series of sequential pre-enrichment culturing steps are needed.  The methods include pre-enrichment on selective or differential media for 16 to 24 hours to increase the levels of Salmonella organisms.

PRINCIPLE OF THE TEST

NHD SMART II Salmonella Test is designed to detect Salmonella antigens in contaminated food samples.  The test is a rapid qualitative test that is based on the use of Salmonella-specific antibodies.  The reaction between the enriched positive sample and the colored particle-conjugated antibody will form a complex that migrates along the test membrane.  An immobilized capture antibody will form a red line at the test line (T) area upon reacting with the sample/gold complex.  An internal control line (C) is included with the test cassette to assure that the test has been processed correctly.TEST PROCEDURE

Materials Provided

- NHD SMART II Salmonella test cassettes

- Salmonella Sample Diluent Buffer  1 bottle

- Droppers

- Positive Control

Materials NOT Provided

Enrichment media – We recommend Selenite Cysteine Broth (GN Broth - Hanja) or (Difco, Cat. # 0486-05 or Accumedia, Cat. # 7218).  Alternatively, Selenite Cystine Broth (Difco, Cat. # 0687-05 or Accumedia Cat. # 7283) may also be used.

Also required are the necessary ancillary equipment, such as stomacher bag, test tubes, pipettes, and safety equipment used in a microbiology testing laboratory for preparation, storage, and handling of samples.

Sample Collection and Handling

Only food samples should be tested.

Samples that will not be tested within 48 hours should be refrigerated at 2-8°C.

Protocol

Allow the pre-enriched samples, controls, and buffers to reach room temperature, 20-25ºC, prior to use.

Sample Preparation/Test Procedure

1.	A ratio of 1:10 sample to enrichment media (EM) is recommended.  For example, 10 g of ground meat in 90 mL of EM or 25 mL of juice in 225 mL of EM.  Incubate the sample/EM broth mixture at 37ºC for 16-24 hours.

2.	Add 100 µL of Sample Diluent Buffer solution (1X PBS) to a test tube.

3.	Add 100 µL of the enriched sample (from Step 1) to the test tube.  Shake well to mix.

4.	Add 100 µL of the buffer/sample mixture (from Step 3) to the sample well of the test cassette.

5.	Read results within 20-25 minutes and no later than 30 minutes.  Some positive results may be observed within a few minutes depending on the concentration of the antigen in the sample tested.

QUALITY CONTROL

Each test cassette has a built in procedural control.  Correct procedural technique and test device performance is confirmed when a colored line appears in the control (C) area of the test cassette.  It is recommended that when a new shipment of product is received, negative and positive controls be tested, and the appropriate results obtained.

INTERPRETATION OF RESULTS

A reddish colored line developing in the control line area (C) of the test cassette indicates proper sample migration, capillary flow, and reagent integrity.  If this line does not develop, the test is INVALID and must be repeated.

A reddish colored line in the test line area (T) of the test cassette is indicative of a POSITIVE result.  No discernible development of a reddish colored line is indicative of a NEGATIVE result.

Report as follows:

POSITIVE =	Sample Positive for detection of Salmonella.

NEGATIVE =	Sample absent or below the level of analytical sensitivity for Salmonella.

Depending on the concentration of antigen and consistency of sample, POSITIVE results may be observed in as few as two (2) minutes; however, to confirm NEGATIVE results, the complete reaction time of thirty (30) minutes is required.

If test results are indeterminate (very faint reddish color on the test line [T]), the test should be repeated or another sample collected.  A repeatedly faint line developing on the test line of the cassette should be considered a POSITIVE test result.

STORAGE

Store all components of the SMART II Salmonella Test Kit at room temperature (20-25°C).  The test cassettes must remain sealed in the desiccated pouch until ready for use.

LIMITATIONS OF THE PROCEDURE

This test is designed as an aid in the detection of Salmonella and should be used as an adjunct to culture.  Samples containing large amounts of food matter may migrate slowly up the test cassette membrane.  Follow the instructions recommended in the test procedure for transferring these samples to a test tube; and allow a maximum of 30 minutes for migration to occur.  On occasion, samples containing very high concentrations of Salmonella lipopolysaccharide antigens may produce a strong POSITIVE test line (T) and a weak or faint CONTROL line (C) on the test cassette.  This is due to large amounts of conjugate being deposited on the TEST line (T), depleting the amount of conjugate available for binding at the CONTROL line (C).  If a strong positive sample is suspected, the sample may be serially diluted in distilled water to view the CONTROL line (C) and ensure proper performance of the test procedure.

REAGENTS

Test Cassettes:

Immunochromatographic assay cassettes containing dried colloidal gold labeled anti-Salmonella antibody and immobilized polyclonal anti-Salmonella antibody.

Sample Diluent Buffer:

1X Phosphate Buffered Saline (PBS)

PRECAUTIONS

SMART II Salmonella Test Kit is for Food Testing Only.  All samples and test materials should be handled as if capable of transmitting disease.  Strict adherence to established micro-biological safety precautions should be observed when handling samples, test materials and in the disposal of processed assay materials.  Do not use reagents beyond the indicated expiration date.

.